Soft Glue Logic to Control the Tuning Fork Shutter at 8id

By: John Weizeorick

Date: 4-18-2012

The tuning fork shutter runs at a base frequency of 50Hz where the shutter is open for 10msec and closed 10msec.  The shutter can be synchronized with an external clock running at near 50Hz.

The Soft Glue is used to generate this 50Hz clock to the tuning fork shutter as well as providing external triggers to the 480 FCCD camera.  In this way the exposures will be in sync with the shutter.

There are a number of process variables used to control the process of generating a sequence of exposures.

- Start_In  - A Bi which starts a sequence of exposures when toggled high and low

- Stop_In  - A Bi which is used to abort a sequence of exposures

- NumImages – Number of images to take after Start_In has been toggled

There are a number of output signals which should be connected to the appropriate hardware.

- Tuning_Fork_Shutter_Input – 50Hz clock with a 50% duty cycle which goes into the tuning fork input.

- ExtTrg_To_FCCD – External triggers that go into the 480 FCCD camera

- OldSutter_Input – Goes to a slower shutter and opens it during the series of exposures.

The soft glue uses an internal 8MHz base clock, which is divided by 80,000 to generate a pulse every 100Hz.  This pulse is then sent into a D flip-flop which toggles its output states every time there is a pulse.  So the output of this D flip-flop provides a 50Hz clock with a 50% duty cycle.  This clock is used by the tuning fork shutter to synchronize it’s open/close states.

Before a sequence of exposures is taken a user or SPEC script must set the number of images to be collected by writing a value to the NumImages PV in the glue logic.  This number should match the PV on the detector IOC and the PV on the image capture IOC.  To take a series of images the user or the SPEC script needs to toggle the Start_In pv. This will initiate a series of trigger pulses that are in sync with the shutter.  By monitoring the OldSutter_Input the SPEC script will know when the series of images is complete.

There are a number of detector PVs that must be set properly to match the 50Hz cycle time and 10msec exposure time.

The logic used in the soft glue is shown in the schematic below.